INSTRUCTION MANUAL

Duct type oxygen sensor

JXBS-3005-O2

VER1.0

WeihaiJXCTElectronicTechnologyCo.,Ltd.

Ⅰ Brief Introduction

1.1 Product Overview

  - Ducted oxygen sensors are designed specifically for plug-in pipe gas measurement. The plug-in structure with stainless steel gas guide pipe combined with stainless steel flange is adopted as the core detection device, which has the characteristics of wide measuring range, high precision, good linearity, good generality, convenient use, easy installation, long transmission distance and moderate price.

1.2 Primary Parameters

| PARAMETERS | TECHNICAL SPECIFICATIONS |
| --- | --- |
| MEASURING RANGE | 0-30%/0-100% |
| MEASURING MODE | Electrochemical sensor |
| PRECISION | 3%F.s |
| RESPONSE TIME | less than 15 seconds |
| WARRANTY PERIOD | 2 years ( Host ) / 1 year ( Sensor ) |
| BAND RATE | 2400/4800/9600 |
| COMMUNICATION PORT | Analog (voltage / current) / digital (RS485) |
| POWER SUPPLY | 12-24V DC |
| POWER CONSUMPTION | ≤0.15W（@12V DC , 25℃） |
| OPERATING TEMPERATURE | -30-50 ℃（-20-40 ℃ continued） |
| WORKING HUMIDITY ENVIRONMENT | 15-95%RH |
| DUCT LENGTH | 200mm |
| CASE SIZE | 110 x 85 x 44mm3 |

1.3 Probe Parameters And Selection

| NO. | MANUFACTURER | RANGE | RESOLUTION | LIFE |
| --- | --- | --- | --- | --- |
| 30A | British Imports of CITY | 30% | 0.1% | >2 years |
| 30L | Long life of Japanese imports | 100% | 0.1% | >2 years |

1.4 Analog communication parameters

| Parameter | Index |
| --- | --- |
| Current output type | 4-20mA |
| Current output load | ≤600Ω |
| Voltage output type | 0-5V/0-10V |
| Voltage output load | ≤250Ω |

1.5 Digital communication parameters

| Parameter | Index |
| --- | --- |
| Communication Interface | RS485 |
| Baud rate | 2400/4800/9600 |
| Protocol | Modbus RTU |

Ⅱ Hardware Connections

2.1 CHECKING BEFORE INSTALLATION

Check the list of devices before installation:

TABLE 3   List of Devices

| Name | Number |
| --- | --- |
| THE SENSOR DEVICE | 1 |
| 12V POWER ADAPTER（Optional） | 1 |
| THE USB TO 485 DEVICE（Optional） | 1 |
| WARRANTY CARD / CERTIFICATE | 1 |

2.2 Interface Description

The power interface is wide voltage power input 12-24V. When wiring the 485 signal line, pay attention that the two A / B lines cannot be reversed, and the addresses between multiple devices on the bus must not conflict.

|  | Line Color |  | Description |
| --- | --- | --- | --- |
| Power | Brown |  | Power supply Positive ( 12-24V DC ) |
|  | Black |  | Power supply Negative |
| Communication | Yellow ( Gray ) |  | 485-A |
|  | Blue |  | 485-B |

The power interface is wide voltage power input 12-24V. For analog products, pay attention to the positive and negative signal lines. Do not reverse the positive / negative current / voltage signal lines.

|  | Line Color |  | Description |
| --- | --- | --- | --- |
| Power | Brown |  | Power supply Positive ( 12-24V DC ) |
|  | Black |  | Power supply Negative |
| Communication | Yellow ( Gray ) |  | Positive voltage / current output |
|  | Blue |  | Voltage / current output negative |

We provide default cable length of 0.6 meters, you can extend the cable yourself according to your needs.

2.3 Installation Description

The diameter of the duct of the duct sensor is 25mm. The duct sensor is delivered with a set of mounting flanges for fixing to the designated equipment pipe wall. The flanges are fixed by three 3.5mm mounting screw holes. Customers can use 3mm bolts or 3mm self-tapping screws to install according to the situation

There are screws with a distance of 105mm on the sensor housing, which can also be fixed by the fixing screws here.

The following figure shows the installation diagram of the flange. After fixing the flange, the customer can use the screws on the flange post to fix the sensor.

The following figure shows the installed reference style. The specific installation forms are various. Customers can choose a reasonable installation method according to the site conditions:

Ⅲ Digital sensor communication protocol

3.1 Communication Basic Parameters

| PARAMETERS | CONTENT |
| --- | --- |
| Protocol | Modbus RTU |
| Data bits | 8 bit |
| Parity bit | No |
| Stop bit | 1 bit |
| Error checking | CRC (redundant loop code) |
| Baud rate | 2400 bps/ 4800 bps/ 9600 bps can be set factory defaults to 9600 bps |

3.2 Data frame format definition

Modbus-RTU communication protocol is adopted, the format is as follows:

Time for initial structure ≥ 4 bytes

Address code = 1 byte

Function code = 1 byte

Data area = N bytes

Error check = 16-bit CRC

Time to end structure ≥ 4 bytes

Address code: It is the address of the transmitter and is unique in the communication network (factory default 0x01).

Function code: The function instruction of the command issued by the host, this transmitter only uses the function code 0x03 (reading register data).

Data area: The data area is the specific communication data. Note that the high byte of the 16bits data comes first!

CRC code: two-byte check code.

Inquiry frame

| address code | function code | Register start address | Register length | Low check bit | Check code high |
| --- | --- | --- | --- | --- | --- |
| 1 byte | 1 byte | 2 bytes | 2 bytes | 1 byte | 1 byte |

Response frame

| address code | function code | Number of valid bytes | Data area | Second data area | Nth data area |
| --- | --- | --- | --- | --- | --- |
| 1 byte | 1 byte | 1 byte | 2 bytes | 2 bytes | 2 bytes |

3.3 Register Address

| Register Address | Plc Configuration Address | Content | Operation |
| --- | --- | --- | --- |
| 0000H | 40001 | Humidity ( Unit 0.1%RH) | Read-Only |
| 0001H | 40002 | Temperature ( Unit 0.1℃) | Read-Only |
| 0006H | 40007 | Oxygen ( unit 0.1%) | Read-Only |
| 0100H | 40101 | Device Address (0-252) | R/W |
| 0101H | 40102 | Baud Rate (2400/4800/9600) | R/W |

3.4 Communication protocol examples and explanations

3.4.1 Read Device Address 0x01's Oxygen Concentration

Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00 0x06 | 0x00 0x01 | 0x64 | 0x0B |

Answer Frames( For example, the reading is 18.9% )

| Address Code | Function Code | Returns to The Number Of Valid Bytes | Oxygen Value | Check Digit Low | Check Digit High |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x02 | 0x00 0xBD | 0x78 | 0x35 |

Oxygen：

0xBD H ( hexadecimal ) =189=>Oxygen=18.9%

3.4.2 Read Device Address 0x01's Temperature And Humidity Value

Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00,0x00 | 0x00,0x02 | 0xC4 | 0x0B |

Answer Frame

| Address Code | Function Code | Number Of Valid Bytes | Humidity Value | Temperature Value | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x04 | 0x00 0xFE | 0x00 0xAF | 0xDB | 0xBF |

Temperature：

00AF H ( hexadecimal ) =175=> Temperature =17.5 ℃

Humidity：

00FE H ( hexadecimal ) =254=> Humidity =25.4%RH

3.4.3 Read Device Address 0x01's Temperature And Humidity, Oxygen Concentration Value

Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00,0x00 | 0x00,0x07 | 0x04 | 0x08 |

Answer Frame

| Address Code | Function Code | Number Of Valid Bytes | Humidity Value | Temperature Value |
| --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x0E | 0x03 0x14 | 0x01 0x1B |
| 8 Useless Bytes |  | Oxygen Value | CRC_L | CRC_H |
| 0x00 ... |  | 0x00 0x28 | 0x50 | 0x3B |

Temperature：

011B H ( hexadecimal ) =283=> Temperature =28.3 ℃

Humidity：

0314 H ( hexadecimal ) =788=> Humidity =78.8%RH

Oxygen ：

0028 H ( hexadecimal ) =40=> Oxygen  =4%

IV Meaning and Conversion of Analog Sensor Parameters

4.1 Analog 4-20mA Current Output

| Current Value | O3 |
| --- | --- |
| 4mA | 0ppm |
| 20mA | 100ppm |

The formula isP(O2) = (I (current) -4mA) *1.875%

where the unit of P is% and the unit of I is mA.

For example, the data collected in the current situation is 15.15 mA, and the value of O2 is calculated at 20.9%.

The above is a 30% range calculation method, other ranges should be 4 mA for 0%,20mA for the maximum range linear conversion.

4.2 Analogue 0-10V Voltage Output

| Voltage Value | O2 |
| --- | --- |
| 0V | 0% |
| 10V | 30% |

The formula is P(O2)= V (voltage) *0.003%

The unit of P is% and the unit of V is mV.

For example, the data collected in the current case is 6966 mV, at which time the value of the calculated O2 is 20.9%.

The above is a 30% range calculation method, other ranges should be 0 V for 0%,10 V for the maximum range of linear conversion.

4.3 Analogue 0-5V voltage output

| Voltage Value | O2 |
| --- | --- |
| 0V | 0% |
| 10V | 30% |

The formula is P(O2) =V(voltage)*0.006%

The unit of P is % and the unit of V is mV.

For example, the data collected in the current situation is 3483mV, and the calculated value of O2 is 20.9%.

V Appendix

5.1 Additional Product Instructions

《How to Use Multimeter-Assisted Debugging》: Multimeter is a very important auxiliary debugging tool, and it is necessary to use multimeter-Assisted debugging whenever you find errors in readings and your own expectations.

《Quantification Precision and Resolution Description of Analog Products》: The concept of quantization accuracy and resolution of analog products, how to calculate, how to select the range correctly.

《The source and elimination of the deviation of the analog product》: The cause of the error, the type of error and how the error is eliminated in the analog product when there is an error in the reading.

5.2 Warranty and aftermarket

The warranty terms are in accordance with the sensor after-sales clauses of Weihai Jingxun Changtong Electronic Technology Co., Ltd. The warranty period is two years for the main circuit of the sensor, one year for the gas-sensitive probe, and three months for the accessories (shell, plug, cable, etc.)