- Product overview

This product is used to measure the concentration of hydrogen cyanide gas. It adopts a three-electrode structure design and uses an electrochemical method to measure the gas concentration.

- Performan        ce parameter

·Working conditions

  - Physical property

- Model	               JEC-HCN-50

- Normal detection range 	      50ppm

- Maximum detection concentration	100ppm

- Sensitivity                         0.1±0.02uA/ppm

- Response time(T90)                 ≤120 seconds

- Baseline drift       	               -1-0ppm

- Resolution  	               0.2ppm

- Linearity 	               linearity

- Baseline	               ＜±0.2uA

- Bias voltage 	                0mV

- Temperature range              0~50℃

- Humidity range                 0~90%RH (non-condensing)

- Pressure range                 1±0.1 (standard atmospheric pressure)

- Service life                   3 years (in air)

- Long-term stability            <2% signal value/month

- Storage temperature            10~30℃

- Storage time                   6 months (special packaging box)

- Shelf life                     24 months

- Shell material   ABS

- Weight           5g

- Directionality  no

·Product size

※All dimensions are in millimeters unless otherwise specified, all tolerances are ±0.10 mm

·Installation instructions

Please connect the electrode pins correctly. The wrong connection position will affect the normal operation of the sensor; the sensor should avoid contact with organic solvents, paints, oils and high-concentration gases; do not use adhesives directly on the sensor or near the sensor. Prevent the plastic shell from cracking; if it is used in a special environment, please contact us in advance.

| - Cross sensitivity |  |  |
| --- | --- | --- |
| gas | thickness（ppm） | Output signal (equivalent to ppm hydrogen cyanide) |
| carbon monoxide | 5 | 1.5 |
| sulfur dioxide | 15 | 30 |
| nitrogen dioxide | 35 | -1 |
| hepatic gas | 5 | -3 |

  - Temperature effect

    - Use instructions

- 1. All the above performance specifications are measured under environmental conditions: temperature 20℃, relative humidity 50%, and one atmosphere (100kPa).

- 2. It is recommended to use target gas for calibration. If cross-sensitive gas is used for calibration, the accuracy of calibration and measurement cannot be guaranteed.

- 3. The cross sensitivity will fluctuate by ±30%, and it may vary with the production batch of the sensor and the life of the sensor.

- 4. The aforementioned cross-sensitivity includes but is not limited to the aforementioned gases, and the sensor may respond to other gases.